## Monthly QA Sheet
Proton Monthly QA
Rutherford Cancer Centres
v3.1 10/05/2022
Machine
Save measurement data from the MatriXX, Zebra and XRV ect. in the folder for this months QA
(sections in white must be filled in manually)
1. Absolute dosimetry
Setup: | PPC05 RW3 plate phantom, Chamber at 2 cm depth (17 mm RW3 + RW3 cap on PPC05) Chamber effective point of measurement at isocentre (2.5 mm below chamber holder plate surface)
Beam: | 10x10 single energy layers, 0.5 MU per spot, no range shifter
Action level (%): | 2
Tolerance (%): | 3
Chamber |  |  | Bias Voltage (V) |  | #N/A
Electrometer |  |  | ND,w,Q0 (Gy/C): |  | #N/A
Phantom Temperature (°C) |  |  | Kpol×Ks: |  | #N/A
Pressure (hPa) |  |  | KQ,Q0: |  | #N/A
Ktp |  | 1 | Kelec: |  | #N/A
Energy (MeV) |  | 200 | 180 | 120 | 100 | 80
Open Beam Range (g/cm2) |  | #N/A | #N/A | #N/A | #N/A | #N/A
MU delivered per layer |  | #N/A | #N/A | #N/A | #N/A | #N/A
Measurement 1 (nC)
Measurement 2 (nC)
Measurement 3 (nC)
Average (nC)
Measured dose (cGy)
baseline dose (cGy) |  | #N/A | #N/A | #N/A | #N/A | #N/A
Difference from baseline (%)
Max diff (%): |  |  |  |  | Performed by
Result: | Incomplete |  |  |  | Physicist | Date
2. Dose-MU linearity
Setup: | As above or below i.e. PPC05 in RW3 or attached to the centre of the range shifter
Beam: | 5x5 or 10x10 single energy layer
Action level (%): | 1
Tolerance (%): | 2
energy (MeV)
gantry (°)
MU/spot | Beam MU | Charge (nC) | Difference from 1 MU/spot (%) | Result
0.02 | 100.86 |  |  |  | (Action level = tolerance level for 0.02MU spot)
0.05 | 252.15
0.1 | 504.3
0.5 | 2521.5
1 | 5043
5 | 25215 |  |  |  | Performed by
[5 empty] | Physicist | Date
3. Output dependence on gantry angle
Setup: | PPC05 attached to the centre of the range shifter (fully retracted)
Beam: | 10x10 single energy layer
Energy (MeV):
Action level (%): | 1.5
Tolerance (%): | 2
Gantry angle (°) | Charge (nC) or Dose (CGE) | Deviation (%)
0
180
120
90
30
330 |  |  |  |  | Performed by
Max deviation (%): |  |  |  |  | Physicist | Date
Result: | Incomplete
4. Range of single energy layers
Setup: | Zebra aligned with lasers (Zebra surface at isocentre), gantry 90°
Beam: | 10x10 Layers
Action level (cm): | 0.1
Tolerance (cm): | 0.15
Beam energy (MeV) |  | Range shown in TCS | Expected (cm) | Measured range R80 (cm) | Result
226.2 + RS4cm |  | #N/A | #N/A
220 + RS4cm |  | #N/A | #N/A
200 + RS4cm |  | #N/A | #N/A
180 + RS4cm |  | #N/A | #N/A
200 |  | #N/A | #N/A
180 |  | #N/A | #N/A
160 |  | #N/A | #N/A
140 |  | #N/A | #N/A
120 |  | #N/A | #N/A
100 |  | #N/A | #N/A
70 |  | #N/A | #N/A
Performed by
[5 empty] | Physicist | Date
5. Range and SOBP consistancy
Setup: | Zebra aligned with lasers (Zebra surface at isocentre), gantry 90° Move PPS laterally such that the isocenter is at the centre of the SOBP (iso depths in table below). Where a range shifter is needed (given in table below) set the snout position to that prescribed for the field. Export the measurements from OmniPro Incline and rename as described in table below. Copy all the measurements to a folder ..\Analysis software\octave_SOBP_depth_QA\yyyymmdd Run the octave script and copy the results into the table below.
Beam: | SOBPs
| Action levels | Tolerances
R90 (cm): | 0.1 | 0.15
Modulation (cm): | 0.3 | 0.5
Flatness (%): | 2 | 3
Distal 80-20 (cm): | 0.1 | 0.2
Beam |  | Test | Expected | Measured | Result
High energy beam |  | R90 (cm) | #N/A
2010W |  | Modulation (cm) | #N/A
Save as: R20M10n.csv |  | Flatness (%) | #N/A
iso at 15.0 cm depth, RS out |  | Distal 80-20 (cm) | #N/A
Medium energy beam |  | R90 (cm) | #N/A
1006A |  | Modulation (cm) | #N/A
Save as: R10M06n.csv |  | Flatness (%) | #N/A
iso at 7.0 cm depth, RS in |  | Distal 80-20 (cm) | #N/A
Medium energy beam |  | R90 (cm) | #N/A
1005A |  | Modulation (cm) | #N/A
Save as: R10M05n.csv |  | Flatness (%) | #N/A
iso at 7.5 cm depth, RS out |  | Distal 80-20 (cm) | #N/A
Low energy beam |  | R90 (cm) | #N/A
1010A |  | Modulation (cm) | #N/A
Save as: R10M10n.csv |  | Flatness (%) | #N/A
iso at 5.0 cm depth, RS in |  | Distal 80-20 (cm) | #N/A
Performed by
[5 empty] | Physicist | Date
6. Spot size, shape and relative spot position accuracy
Setup: | Lynx, aligned to isocentre with lasers at saved couch position with required pitch to remove any rotation of the image, tail torwards T In "..\Analysis software\SpotPositionAnalysis", create opg with "dicom2opg", copy opg files to "Spot_pattern_opg_files_to_analyse", analyse with "Run Spotemkine and Analysis.bat" octave script
Beam: | Spot size, 17 spots, no range shifter
Gantry angle:
|  | Action levels | Tolerances
Sigma (%): |  | 10 | 15
Skewness: |  | 0.2 | 0.5
Position (No. points > 1mm off): |  | 1 | 2
max. dev. (mm): |  | 1.5 | 2
Energy (MeV) | SigmaX (mm) | SigmaY (mm) | SkewX | SkewY | No. of points > 1mm off | max. dev. (mm) | Result
226.2
Diff to baseline (%):
200
Diff to baseline (%):
150
Diff to baseline (%):
100
Diff to baseline (%):
80
Diff to baseline (%):
70
Diff to baseline (%):
Performed by
[5 empty] | Physicist | Date
7. Lateral uniformity
Setup: | MatriXX on PPS at isocentre with no buildup Align MatriXX against a lokbar on H4, GOTO Treatment PPS 1. Measure in MyQA FastTrack, use "VARIAN" photon analysis: use 40 seconds single shot for 226 and 150 MeV and 70s for 70MeV Record the results from Fast track in the table below Save the .opg files to: ..\Monthly\20XX-XX-XX\MatriXX - Lateral Uniformity 2. Open the "MatriXX Lat Unif Analysis" excel file, import the opg files, analyse and enter the max deviation for the 3 energies below
Beam: | 20x20 single energy layers, no range shifter
| Action levels | Tolerances
Flatness and Symmetry (%) | 3 | 4
Max chamber deviation (%) | 4 | 5
Field Size (mm) | 2 | 3
Penumbra (mm) | 2 | 3
Beam |  | Test | Expected | Measured | Result
High energy beam (MeV) |  | Max Flatness (%) | #N/A
226.2 |  | Max Symmetry (%) | #N/A
#N/A |  | Field Size Inline (mm) | #N/A
|  | Field Size Crossline (mm) | #N/A
|  | Max Penumbra 80-20 (mm) | #N/A
Medium energy beam (MeV) |  | Max Flatness (%) | #N/A
150 |  | Max Symmetry (%) | #N/A
#N/A |  | Field Size Inline (mm) | #N/A
|  | Field Size Crossline (mm) | #N/A
|  | Max Penumbra 80-20 (mm) | #N/A
Low energy beam (MeV) |  | Max Flatness (%) | #N/A
70 |  | Max Symmetry (%) | #N/A
#N/A |  | Field Size Inline (mm) | #N/A
|  | Field Size Crossline (mm) | #N/A
|  | Max Penumbra 80-20 (mm) | #N/A
Beam energy (MeV) |  | Max. Single Chamber Deviation from Av. (%) | Result
226.2 |  |  |  |  | Performed by
150 |  |  |  |  | Physicist | Date
70
8. Proton beam to X-ray isocentre alignment
Setup: | XRV-100 aligned to the lasers on the fourth spot from the tip with the PPS in the saved XRV position XRV-100 workflow in Adapt Insight Physics mode. Aligned using obliques and the 8 markers BeamWorks - General Isocentre(Automatic) script on XRV-100 laptop Analyse output.txt with: S:\Clinical\Radiotherapy\Physics\QA\NP10\Proton\SpotPositionAnalysis\XRV-100 Analysis.xlsx
Beam: | XRV-100 single spot at isocentre, no range shifter
| Action levels | Tolerances
Position (mm): | 1 | 2
Angle (°): | 0.5 | 1
Position of the marker relative to the proton spot
Energies (MeV) |  | 226.2, 200, 160, 100 & 70 | Result
Requested Gantry Angle (°)
Max gantry angle error (°)
Max gantry tilt (°)
Max beam distance to iso (mm)
Requested Gantry Angle (°)
Max gantry angle error (°)
Max gantry tilt (°)
Max beam distance to iso (mm)
Requested Gantry Angle (°)
Max gantry angle error (°)
Max gantry tilt (°)
Max beam distance to iso (mm)
Requested Gantry Angle (°)
Max gantry angle error (°) |  |  |  |  | Performed by
Max gantry tilt (°) |  |  |  |  | Physicist | Date
Max beam distance to iso (mm)
9. Isocentric PPS rotation
Setup: | Marker stuck on PPS (or phantom) anywhere superior of H2, aligned manually to isocentre with lasers (PPS at zero rotation), couch not loaded with any weight. Load all fields, go to setup field, align marker to isocenter with obliques using Sphinx workspace in Adapt Insight Physics mode (PPS at zero rotation), apply correction, confirm position, go to next field and rotate PPS. For each PPS rotation align the marker with the obliques and record the correction but do not apply it.
Beam: | Couch rotation plan with a site setup and fields at each PPS angle
Action level (cm): | 0.1
Tolerance (cm): | 0.15
PPS angle (°) | Correction Vector |  |  | Position error
| x (cm) | y (cm) | z (cm) | (cm)
315
270
225
180
Max error (cm): |  |  |  |  | Performed by
Result: | Incomplete |  |  |  | Physicist | Date
10. kV 3D Image Quality
Setup: | Setup CatPhan with lasers to the 2nd alignment marker from the G-end of the CatPhan. In Adapt Insight use "QA CBCT CatPhan" workflow in physics mode. Aqcuire CBCT image using the presets given in the table below. Score the images for spatial resolution in Adapt Insight then export the image to be analysed in IQWorks (contrast, uniformity) and Sante DICOM Viewer (geometric scale). Review a prostate patient CBCT from this month, or if there are no suitable patients scan the large water bottle with the Pelvis or Prostate clinical preset
Action levels: | Uniformity (%): |  | 4 | Contrast (%): |  | 3
| Resolution (min line pair groups): |  | 7 | Scale (Vert, Horiz) (mm): |  | 1 | Scale (Sag) (mm): | 2
| Uniformity |  | Contrast |  |  | 3D Spatial Resolution | Geometric Scale (mm) |  | Result
|  |  | Insert/Parameter | Value | Low Contrast Visibility | Visible Line Pairs
Preset: Catphan LFOV | Centre |  | Polystyrene (7oclock) Mean |  |  |  | Vert. | Diff to Expected
| North |  | Polystyrene (7oclock) SD |  |  |  | Horiz. | Diff to Expected
| South |  | LDPE (9 o'clock) Mean |  |  |  | Sag. | Diff to Expected
| East |  | LDPE (9 o'clock) SD
West
Difference (%)
Preset: Catphan SFOV | Centre |  | Polystyrene (7oclock) Mean |  |  |  | Vert. | Diff to Expected
| North |  | Polystyrene (7oclock) SD |  |  |  | Horiz. | Diff to Expected
| South |  | LDPE (9 o'clock) Mean |  |  |  | Sag. | Diff to Expected
| East |  | LDPE (9 o'clock) SD
West
Difference (%)
Qualitative review of a LFOV clinical preset CBCT of a patient or large water bottle
Preset: |  | Image quality acceptable for matching?
Performed by
[5 empty] | Physicist | Date
11. kV 2D Image Quality
Setup: | Leeds test object taped to detector panel. Copper plate placed over corresponding tube window on floor. In Adapt Insight use "kV/kV Oblique XRV-100" workspace in physics mode. Aqcuire image using "Pelvis" preset then repeat for the other tube/detector. Score the images in Adapt Insight.
Action levels: | Contrast (min disks visible): |  | 12 | Resolution (min group resolvable): |  | 10
| Low Contrast Visibility | 2D Spatial Resolution | Pass/Fail
OB1 (A-side detector)
OB2 (B-side detector) |  |  |  |  | Performed by
[5 empty] | Physicist | Date
12. kV Output
Setup: | Tape a Farmer chamber (in protective sheeth) to the PPS such that it overhangs the superior end. Set chamber centre to isocentre using lasers. Electrometer set to low range. In Adapt Insight use "QA CBCT CatPhan" and "kV/kV Oblique XRV-100" workspaces in physics mode to perform the image acquisitions given in the table below. For the oblique acquisitions ensure only the tube being measured is selected for exposure.
Action level (%): | 10
Baseline chamber SN: |  | #N/A
Measurement Chamber SN:
Temperature (°):
Pressure (hPa):
Image Acquisition Type/Preset |  | Charge Obtained (pC) | TP corrected charge (pC) | TP corrected baseline (pC) | Diff to Baseline (%) | Result
CBCT "Head" |  |  |  | #N/A
CBCT "Pelvis" |  |  |  | #N/A
Oblique OB1 "Pelvis" |  |  |  | #N/A
Oblique OB2 "Pelvis" |  |  |  | #N/A
Performed by
[5 empty] | Physicist | Date
13. Check of interlocks and warning lights
Setup: | Check that the interlocks stop the beam, and the warning lights, laser curtains and CBCT light guard are operational and the PPS collision detection works as intended
Maze entrance light curtain
Maze motion sensor and CBCT light guard
PPS collision detection
TCS pause
DCEU pause |  |  |  |  | Performed by
Beam on and X-Ray on, outside room |  |  |  |  | Physicist | Date
Traffic lights
14. Daily QA review. |  |  |  |  | Performed by
Daily QA Passed everyday in previous month: |  |  |  |  | Physicist | Date
All failures have been investigated and solved:
15. Monthly QA Review comments, or comments on failures.
Comments and followup of failures |  |  |  |  | Physicist | Date
Note on C43: Edward Ilsley: MU to be entered into Required MU in Beam Selection.
Note on B192: Edward Ilsley: "Vector" as given in XRV-100 analysis spreadsheet.
Note on B196: Edward Ilsley: "Vector" as given in XRV-100 analysis spreadsheet.
Note on B200: Edward Ilsley: "Vector" as given in XRV-100 analysis spreadsheet.
Note on B204: Edward Ilsley: "Vector" as given in XRV-100 analysis spreadsheet.
Note on B281: Edward Ilsley: Break the maze entrance light curtain while a proton beam is being delivered and ensure the proton beam switches off.
Note on B282: Edward Ilsley: Ensure the LEDs on the light guard in the ceiling half way through the maze activates when motion under the light guard. Ensure to check it activates with movement along walls of the maze rather than just the centre of the maze. Ensure the CBCT light guard activates while the panel and tube are being extended or retracted by breaking the light curtain on both tube side and panel side individually.
Note on B283: Edward Ilsley: Manually move the PPS with the hand pendant and ensure that the collision interlock activates when physical resistence is added to the couch movement.
Note on B284: Edward Ilsley: Ensure the beam terminates upon pressing "Pause" on Adapt Deliver while a beam is running.
Note on B285: Edward Ilsley: Prepare a beam and set an MU limit on the DCEU that is lower than the total MU for that beam. Run the beam and check the beam terminates at the MU limit set on the DCEU.
Note on B287: Edward Ilsley: Yellow traffic light is on when maze last-man-out is completed. Red and Yellow traffic lights are on when proton beam is being delivered. Red and Yellow traffic lights are on when kV beams are being delivered (Obliques or CBCT).

## Electrometer and Chamber
Values for the currently selected chamber and electrometer:
Chamber | Chamber Type | Serial Number | ND,w,Q0 (Gy/C) | KQ,Q0 | kpol | ks | Bias (V) | P0 (kPa) | T0 (°C) | Calibration Date | Calibration Expiration Date | Comment
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Electrometer | Serial Number | kelec
#N/A | #N/A | #N/A
Ionisation Chamber Calibration Values
Chamber Name | Chamber Type | Serial Number | ND,w,Q0 (Gy/C) | KQ,Q0 | kpol | ks | Bias (V) | P0 (kPa) | T0 (°C) | Calibration Date | Calibration Expiration Date | Comments
PPC05 SN1109 | Parallel Plate | 1109 | 582266666.6667 | 1.002 | 1 | 1.003 | -400 | 101.325 | 20 | 2021-01-22 | 2023-01-22 | Average of NPL and IBA calibrations
PPC05 SN1110 | Parallel Plate | 1110 | 604000000 | 1.002 | 1 | 1.003 | -400 | 101.325 | 20 | 2018-02-21 | 2022-09-08 | Calibration factor cross checked on 8/9/2020
PPC05 SN1155 | Parallel Plate | 1155 | 569025000 | 1.002 | 1 | 1.003 | -400 | 101.325 | 20 | 2021-01-22 | 2023-01-22 | Average of NPL and IBA calibrations
PPC05 SN1191 | Parallel Plate | 1191 | 597500000 | 1.002 | 1 | 1.003 | -400 | 101.325 | 20 | 2018-09-11 | 2022-09-22 | Calibration factor cross checked on 22/09/2020
PPC05 SN1192 | Parallel Plate | 1192 | 576300000 | 1.002 | 1 | 1.003 | -400 | 101.325 | 20 | 2021-01-22 | 2023-01-22 | Average of NPL and IBA calibrations
PPC05 SN937 | Parallel Plate | 937 | 595500000 | 1.002 | 1 | 1.003 | -400 | 101.325 | 20 | 2016-09-01 | 2022-09-09 | Calibration factor cross checked on 9/9/2020
Reported Electrometer Calibration Values
Electrometer Name | Serial Number | kelec | Comments
Unidos | All Unidos | 1.000 | Any electrometer with a 1.000 factor

## Baseline SOBP
Zebra measurements of SOBPs with the isocentre depth as the physical distance from the front line on the Zebra
Results are from the Octave script "ddcqa9.m"
Baseline values for the currently selected machine (do not modify):
| Lookup index (Monthly QA sheet): | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
|  | Field Name | Patient ID | Site | isocenter depth (cm) | Distal 90% (cm) | Modulation 95-90 (cm) | Flatness (%) | Distal 80-20% (cm) | File name
|  | 1005A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 1006A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 1010A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 2010W | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Individual machine baselines (add new/modify baselines in the table below):
| Lookup index (for above): | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11
Vlookup key | Machine name | Field Name | Patient ID | Site | isocenter depth (cm) | Distal 90% (cm) | Modulation 95-90 (cm) | Flatness (%) | Distal 80-20% (cm) | File name | Date set
iPOne0041005A | iPOne004 | 1005A | 2QA-NP10-241117 | Site 01 | 7.5 | 10.24 | 5.23 | 1.67 | 0.32 | R10M05n.csv | 2022-02-10
iPOne0041006A | iPOne004 | 1006A | 2QA-NP10-241117 | Range Shifter 2nd | 7 | 9.94 | 6.23 | 3.04 | 0.39 | R10M06n.csv | 2022-02-10
iPOne0041010A | iPOne004 | 1010A | 2QA-NP10-241117 | Range Shifter 2nd | 5 | 9.94 | 9.79 | 3.02 | 0.37 | R10M10n.csv | 2022-02-10
iPOne0042010W | iPOne004 | 2010W | 2QA-NP10-241117 | WPE Audit | 15 | 20.04 | 9.5 | 2.18 | 0.51 | R20M10n.csv | 2022-02-10
iPOne0061005A | iPOne006 | 1005A | 2QA-NP10-241117 | Site 01 | 7.5 | 10.15 | 5.2 | 1.89 | 0.34 | R10M05n.csv | 2022-02-10
iPOne0061006A | iPOne006 | 1006A | 2QA-NP10-241117 | Range Shifter 2nd | 7 | 9.93 | 6.2 | 2 | 0.37 | R10M06n.csv | 2022-02-10
iPOne0061010A | iPOne006 | 1010A | 2QA-NP10-241117 | Range Shifter 2nd | 5 | 9.93 | 9.81 | 2.73 | 0.37 | R10M10n.csv | 2022-02-10
iPOne0062010W | iPOne006 | 2010W | 2QA-NP10-241117 | WPE Audit | 15 | 19.97 | 9.4 | 3 | 0.5 | R20M10n.csv | 2022-02-10
iPOne0131005A | iPOne013 | 1005A | 2QA-NP10-241117 | Site 01 | 7.5 | 10.21 | 5.25 | 1.64 | 0.33 | R10M05n.csv | 2022-02-10
iPOne0131006A | iPOne013 | 1006A | 2QA-NP10-241117 | Range Shifter 2nd | 7 | 9.94 | 6.27 | 2.19 | 0.37 | R10M06n.csv | 2022-02-10
iPOne0131010A | iPOne013 | 1010A | 2QA-NP10-241117 | Range Shifter 2nd | 5 | 9.94 | 9.78 | 2.28 | 0.37 | R10M10n.csv | 2022-02-10
iPOne0132010W | iPOne013 | 2010W | 2QA-NP10-241117 | WPE Audit | 15 | 20 | 9.52 | 2.21 | 0.48 | R20M10n.csv | 2022-02-10

## Baseline Monoenergetic
Dose of a single 10x10 cm layer, 2.5 mm spot spacing, 1 MU per spot
2cm depth in water at isocentre
Range and fall off from Zebra measurements
Baseline values for the currently selected machine (do not modify):
| Lookup index (Monthly QA sheet): | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13
[6 empty] | Zebra Range Measurements |  |  |  | MatriXX 20x20 Lateral Uniformity Fields Reference Values
|  | Energy (MeV) | MU | Chamber | Dose (cGy) in water | Distal 90% range (cm) | Distal 80% range (cm) | Distal fall off 80-20 (cm) | "Open beam range" in TCS | Flatness (%) | Symmetry (%) | Field size inline (mm) | Field size crossline (mm) | Penumbra 80-20 (mm)
|  | 226.2 + RS4cm | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 226.09 + RS4cm | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 220 + RS4cm | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 200 + RS4cm | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 180 + RS4cm | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 226.2 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 226.09 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 220 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 215 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 210 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 205 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 200 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 195 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 190 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 185 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 180 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 175 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 170 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 165 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 160 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 155 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 150 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 145 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 140 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 135 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 130 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 125 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 120 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 115 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 110 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 105 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 100 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 95 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 90 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 85 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 80 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 75 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 70.18 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 70 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Individual machine baselines (add new/modify baselines in the table below):
| Lookup index (for above): | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15
[6 empty] | Zebra Range Measurements from 25/06/19 |  |  |  | MatriXX 20x20 Lateral Uniformity Fields Reference Values
Vlookup key | Machine | Energy (MeV) | MU | Chamber | Dose (cGy) in water (23/8/2019) | Distal 90% range (cm) | Distal 80% range (cm) | Distal fall off 80-20 (cm) | "Open beam range" in TCS | Flatness (%) | Symmetry (%) | Field size inline (mm) | Field size crossline (mm) | Penumbra 80-20 (mm)
iPOne013226.2 + RS4cm | iPOne013 | 3: 226.2 + RS4cm; 4: 1681; 5: PPC05 (1110); 7: 28.01; 8: 28.11; 9: 0.5; 10: 32.04
iPOne013220 + RS4cm | iPOne013 | 3: 220 + RS4cm; 4: 1681; 5: PPC05 (1110); 7: 26.53; 8: 26.62; 9: 0.48; 10: 30.56
iPOne013200 + RS4cm | iPOne013 | 3: 200 + RS4cm; 4: 1681; 5: PPC05 (1110); 7: 21.93; 8: 22.02; 9: 0.45; 10: 25.96
iPOne013180 + RS4cm | iPOne013 | 3: 180 + RS4cm; 4: 1681; 5: PPC05 (1110); 7: 17.65; 8: 17.73; 9: 0.41; 10: 21.66
iPOne013226.2 | iPOne013 | 226.2 | 1681 | PPC05 (1110) | 92.4483 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 32.04 | 0 | 0 | 204 | 204 | 9.8
iPOne013220 | iPOne013 | 220 | 1681 | PPC05 (1110) | 92.1152 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 30.56
iPOne013215 | iPOne013 | 215 | 1681 | PPC05 (1110) | 91.9335 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 29.38
iPOne013210 | iPOne013 | 210 | 1681 | PPC05 (1110) | 91.7064 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 28.23
iPOne013205 | iPOne013 | 205 | 1681 | PPC05 (1110) | 91.6156 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 27.09
iPOne013200 | iPOne013 | 200 | 1681 | PPC05 (1110) | 91.4642 | 25.95 | 26.04 | 0.45 | 25.96
iPOne013195 | iPOne013 | 195 | 1681 | PPC05 (1110) | 91.3128 | 24.85 | 24.94 | 0.44 | 24.86
iPOne013190 | iPOne013 | 190 | 1681 | PPC05 (1110) | 91.1008 | 23.75 | 23.84 | 0.43 | 23.77
iPOne013185 | iPOne013 | 185 | 1681 | PPC05 (1110) | 91.1311 | 22.7 | 22.79 | 0.42 | 22.71
iPOne013180 | iPOne013 | 180 | 1681 | PPC05 (1110) | 91.0554 | 21.66 | 21.74 | 0.41 | 21.66
iPOne013175 | iPOne013 | 175 | 1681 | PPC05 (1110) | 90.8737 | 20.6 | 20.69 | 0.4 | 20.63
iPOne013170 | iPOne013 | 170 | 1681 | PPC05 (1110) | 90.8888 | 19.61 | 19.69 | 0.39 | 19.62
iPOne013165 | iPOne013 | 165 | 1681 | PPC05 (1110) | 90.798 | 18.63 | 18.71 | 0.39 | 18.63
iPOne013160 | iPOne013 | 160 | 1681 | PPC05 (1110) | 90.8434 | 17.69 | 17.76 | 0.37 | 17.66
iPOne013155 | iPOne013 | 155 | 1681 | PPC05 (1110) | 90.7677 | 16.69 | 16.77 | 0.37 | 16.71
iPOne013150 | iPOne013 | 150 | 1681 | PPC05 (1110) | 90.8737 | 15.79 | 15.86 | 0.35 | 15.78 | 0 | 0 | 204.5 | 204.5 | 10.3
iPOne013145 | iPOne013 | 145 | 1681 | PPC05 (1110) | 90.9645 | 14.88 | 14.95 | 0.33 | 14.87
iPOne013140 | iPOne013 | 140 | 1681 | PPC05 (1110) | 91.0402 | 13.99 | 14.05 | 0.32 | 13.99
iPOne013135 | iPOne013 | 135 | 1681 | PPC05 (1110) | 91.0857 | 13.12 | 13.18 | 0.31 | 13.12
iPOne013130 | iPOne013 | 130 | 1681 | PPC05 (1110) | 91.2976 | 12.28 | 12.34 | 0.29 | 12.28
iPOne013125 | iPOne013 | 125 | 1681 | PPC05 (1110) | 91.4642 | 11.45 | 11.5 | 0.27 | 11.46
iPOne013120 | iPOne013 | 120 | 1681 | PPC05 (1110) | 91.8578 | 10.68 | 10.73 | 0.26 | 10.66
iPOne013115 | iPOne013 | 115 | 1681 | PPC05 (1110) | 92.2969 | 9.91 | 9.96 | 0.25 | 9.89
iPOne013110 | iPOne013 | 110 | 1681 | PPC05 (1110) | 92.842 | 9.16 | 9.2 | 0.23 | 9.14
iPOne013105 | iPOne013 | 105 | 1681 | PPC05 (1110) | 93.493 | 8.43 | 8.48 | 0.22 | 8.42
iPOne013100 | iPOne013 | 100 | 1681 | PPC05 (1110) | 94.3409 | 7.73 | 7.77 | 0.2 | 7.72
iPOne01395 | iPOne013 | 95 | 1681 | PPC05 (1110) | 95.4159 | 7.06 | 7.1 | 0.18 | 7.05
iPOne01390 | iPOne013 | 90 | 1681 | PPC05 (1110) | 96.7936 | 6.41 | 6.44 | 0.17 | 6.4
iPOne01385 | iPOne013 | 85 | 1681 | PPC05 (1110) | 98.5348 | 5.77 | 5.81 | 0.16 | 5.78
iPOne01380 | iPOne013 | 80 | 1681 | PPC05 (1110) | 100.8816 | 5.18 | 5.21 | 0.15 | 5.18
iPOne01375 | iPOne013 | 75 | 1681 | PPC05 (1110) | 104.1217 | 4.63 | 4.65 | 0.13 | 4.62
iPOne01370 | iPOne013 | 70 | 1681 | PPC05 (1110) | 108.5124 | 4.12 | 4.14 | 0.07 | 4.08 | 0 | 0 | 205 | 205 | 14.1
[6 empty] | Zebra Range Measurements from 16/02/2018 |  |  |  | MatriXX 20x20 Lateral Uniformity Fields Reference Values
MachineEnergy (MeV) | Machine | Energy (MeV) | MU | Chamber | Dose (cGy) in water (08/02/2018) | Distal 90% range (cm) | Distal 80% range (cm) | Distal fall off 80-20 (cm) | "Open beam range" in TCS | Flatness (%) | Symmetry (%) | Field size inline (mm) | Field size crossline (mm) | Penumbra 80-20 (mm)
iPOne004226.2 + RS4cm | iPOne004 | 3: 226.2 + RS4cm; 4: 840.5; 7: 28.01; 8: 28.11; 9: 0.49; 10: 32.04
iPOne004220 + RS4cm | iPOne004 | 3: 220 + RS4cm; 4: 840.5; 7: 26.57; 8: 26.67; 9: 0.49; 10: 30.56
iPOne004200 + RS4cm | iPOne004 | 3: 200 + RS4cm; 4: 840.5; 7: 22.03; 8: 22.12; 9: 0.45; 10: 25.96
iPOne004180 + RS4cm | iPOne004 | 3: 180 + RS4cm; 4: 840.5; 7: 17.71; 8: 17.8; 9: 0.42; 10: 21.66
iPOne004226.2 | iPOne004 | 226.2 | 840.5 | PPC05 (937) | 92.7831 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 32.04 | 0 | 0 | 204 | 204 | 9.8
iPOne004220 | iPOne004 | 220 | 840.5 | PPC05 (937) | 92.5601 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 30.56
iPOne004215 | iPOne004 | 215 | 840.5 | PPC05 (937) | 92.3877 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 29.38
iPOne004210 | iPOne004 | 210 | 840.5 | PPC05 (937) | 92.4587 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 28.23
iPOne004205 | iPOne004 | 205 | 840.5 | PPC05 (937) | 92.266 | Invalid, add RW3 | Invalid, add RW3 | Invalid, add RW3 | 27.09
iPOne004200 | iPOne004 | 200 | 840.5 | PPC05 (937) | 92.1545 | 26 | 26.09 | 0.44 | 25.96
iPOne004195 | iPOne004 | 195 | 840.5 | PPC05 (937) | 92.0936 | 24.92 | 25.01 | 0.44 | 24.86
iPOne004190 | iPOne004 | 190 | 840.5 | PPC05 (937) | 92.1038 | 23.86 | 23.95 | 0.43 | 23.77
iPOne004185 | iPOne004 | 185 | 840.5 | PPC05 (937) | 91.8807 | 22.8 | 22.88 | 0.42 | 22.71
iPOne004180 | iPOne004 | 180 | 840.5 | PPC05 (937) | 91.9821 | 21.77 | 21.85 | 0.43 | 21.66
iPOne004175 | iPOne004 | 175 | 840.5 | PPC05 (937) | 91.83 | 20.71 | 20.79 | 0.42 | 20.63
iPOne004170 | iPOne004 | 170 | 840.5 | PPC05 (937) | 91.7995 | 19.69 | 19.77 | 0.4 | 19.62
iPOne004165 | iPOne004 | 165 | 840.5 | PPC05 (937) | 91.9009 | 18.74 | 18.82 | 0.4 | 18.63
iPOne004160 | iPOne004 | 160 | 840.5 | PPC05 (937) | 91.7894 | 17.74 | 17.82 | 0.39 | 17.66
iPOne004155 | iPOne004 | 155 | 840.5 | PPC05 (937) | 91.7894 | 16.78 | 16.85 | 0.38 | 16.71
iPOne004150 | iPOne004 | 150 | 840.5 | PPC05 (937) | 91.8807 | 15.88 | 15.95 | 0.36 | 15.78 | 0 | 0 | 204.5 | 204.5 | 10.3
iPOne004145 | iPOne004 | 145 | 840.5 | PPC05 (937) | 91.8198 | 14.92 | 14.99 | 0.34 | 14.87
iPOne004140 | iPOne004 | 140 | 840.5 | PPC05 (937) | 91.7083 | 13.99 | 14.06 | 0.32 | 13.99
iPOne004135 | iPOne004 | 135 | 840.5 | PPC05 (937) | 92.0125 | 13.15 | 13.22 | 0.31 | 13.12
iPOne004130 | iPOne004 | 130 | 840.5 | PPC05 (937) | 92.4181 | 12.35 | 12.41 | 0.29 | 12.28
iPOne004125 | iPOne004 | 125 | 840.5 | PPC05 (937) | 92.3471 | 11.47 | 11.53 | 0.28 | 11.46
iPOne004120 | iPOne004 | 120 | 840.5 | PPC05 (937) | 92.7223 | 10.7 | 10.75 | 0.26 | 10.66
iPOne004115 | iPOne004 | 115 | 840.5 | PPC05 (937) | 93.1685 | 9.93 | 9.98 | 0.25 | 9.89
iPOne004110 | iPOne004 | 110 | 840.5 | PPC05 (937) | 93.6045 | 9.18 | 9.23 | 0.23 | 9.14
iPOne004105 | iPOne004 | 105 | 840.5 | PPC05 (937) | 94.2535 | 8.45 | 8.5 | 0.22 | 8.42
iPOne004100 | iPOne004 | 100 | 840.5 | PPC05 (937) | 95.0444 | 7.74 | 7.78 | 0.2 | 7.72
iPOne00495 | iPOne004 | 95 | 840.5 | PPC05 (937) | 95.9976 | 7.06 | 7.09 | 0.18 | 7.05
iPOne00490 | iPOne004 | 90 | 840.5 | PPC05 (937) | 97.2753 | 6.39 | 6.43 | 0.17 | 6.4
iPOne00485 | iPOne004 | 85 | 840.5 | PPC05 (937) | 98.9281 | 5.77 | 5.81 | 0.16 | 5.78
iPOne00480 | iPOne004 | 80 | 840.5 | PPC05 (937) | 101.2705 | 5.19 | 5.22 | 0.14 | 5.18
iPOne00475 | iPOne004 | 75 | 840.5 | PPC05 (937) | 104.3025 | 4.64 | 4.66 | 0.14 | 4.62
iPOne00470 | iPOne004 | 70 | 840.5 | PPC05 (937) | 108.4498 | 4.08 | 4.11 | 0.12 | 4.08 | 0 | 0 | 205 | 205 | 14.1
[6 empty] | Zebra Range Measurements from 17/04/2019 |  |  |  | MatriXX 20x20 Lateral Uniformity Fields Reference Values
MachineEnergy (MeV) | Machine | Energy (MeV) | MU | Chamber | Dose (cGy) in water (09/05/2019) | Distal 90% range (cm) | Distal 80% range (cm) | Distal fall off 80-20 (cm) | "Open beam range" in TCS | Flatness (%) | Symmetry (%) | Field size inline (mm) | Field size crossline (mm) | Penumbra 80-20 (mm)
iPOne006226.09 + RS4cm | iPOne006 | 3: 226.09 + RS4cm; 4: 840.5; 7: 28.04; 8: 28.14; 9: 0.5; 10: 32.01
iPOne006220 + RS4cm | iPOne006 | 3: 220 + RS4cm; 4: 840.5; 7: 26.57; 8: 26.66; 9: 0.49; 10: 30.56
iPOne006200 + RS4cm | iPOne006 | 3: 200 + RS4cm; 4: 840.5; 7: 21.96; 8: 22.05; 9: 0.45; 10: 25.96
iPOne006180 + RS4cm | iPOne006 | 3: 180 + RS4cm; 4: 840.5; 7: 17.69; 8: 17.77; 9: 0.41; 10: 21.66
iPOne006226.09 | iPOne006 | 226.09 | 840.5 | PPC05 (937) | 92.1788 | 31.99 | 32.08 | 0.48 | 32.01 | 0 | 0 | 204 | 204 | 9.8
iPOne006220 | iPOne006 | 220 | 840.5 | PPC05 (937) | 92.1482 | 30.59 | 30.69 | 0.49 | 30.56
iPOne006200 | iPOne006 | 200 | 840.5 | PPC05 (937) | 91.9133 | 25.97 | 26.06 | 0.45 | 25.96
iPOne006180 | iPOne006 | 180 | 840.5 | PPC05 (937) | 91.6683 | 21.68 | 21.76 | 0.42 | 21.66
iPOne006160 | iPOne006 | 160 | 840.5 | PPC05 (937) | 91.1272 | 17.69 | 17.77 | 0.38 | 17.66
iPOne006150 | iPOne006 | 3: 150; 10: 15.78; 11: 0; 12: 0; 13: 204.5; 14: 204.5; 15: 10.3
iPOne006140 | iPOne006 | 140 | 840.5 | PPC05 (937) | 90.7036 | 13.96 | 14.03 | 0.32 | 13.99
iPOne006120 | iPOne006 | 120 | 840.5 | PPC05 (937) | 91.6224 | 10.65 | 10.7 | 0.26 | 10.66
iPOne006100 | iPOne006 | 100 | 840.5 | PPC05 (937) | 93.9246 | 7.73 | 7.77 | 0.2 | 7.72
iPOne00680 | iPOne006 | 80 | 840.5 | PPC05 (937) | 100.0603 | 5.18 | 5.21 | 0.14 | 5.18
iPOne00670.18 | iPOne006 | 70.18 | 840.5 | PPC05 (937) | 106.6656 | 4.1 | 4.13 | 0.12 | 4.1 | 0 | 0 | 205 | 205 | 14.1

## Baseline Spots
Baseline Lynx measurements at isocentre with no build up at gantry 90 deg
Baseline values for the currently selected machine (do not modify):
| Lookup index (Monthly QA sheet): | 1 | 2 | 3 | 4 | 5 | 6
|  | Energy (MeV) | sigmaX (mm) | sigmaY (mm) | skewnessX | skewnessY | Range in TCS
|  | 70 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 70.18 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 75 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 80 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 85 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 90 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 95 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 100 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 105 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 110 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 115 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 120 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 125 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 130 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 135 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 140 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 145 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 150 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 155 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 160 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 165 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 170 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 175 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 180 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 185 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 190 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 195 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 200 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 205 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 210 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 215 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 220 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 226.09 | #N/A | #N/A | #N/A | #N/A | #N/A
|  | 226.2 | #N/A | #N/A | #N/A | #N/A | #N/A
Individual machine baselines (add new/modify baselines in the table below):
| Lookup index (for above): | 3 | 4 | 5 | 6 | 7 | 8
Vlookup Key | Machine | Energy (MeV) | sigmaX (mm) | sigmaY (mm) | skewnessX | skewnessY | Range in TCS
iPOne01370 | iPOne013 | 70 | 7.83 | 7.68 | 0.01 | 0.01 | 4.08 | iPOne013 Baseline measurements from 10/05/2019, analysis with spotemkine
iPOne01375 | iPOne013 | 75 | 7.49 | 7.36 | 0.01 | 0.02 | 4.62
iPOne01380 | iPOne013 | 80 | 7.19 | 7.05 | 0.01 | 0 | 5.18
iPOne01385 | iPOne013 | 85 | 6.85 | 6.7 | 0.01 | 0.02 | 5.78
iPOne01390 | iPOne013 | 90 | 6.52 | 6.37 | 0.03 | 0 | 6.4
iPOne01395 | iPOne013 | 95 | 6.22 | 6.06 | 0 | 0 | 7.05
iPOne013100 | iPOne013 | 100 | 5.92 | 5.77 | 0.01 | 0.01 | 7.72
iPOne013105 | iPOne013 | 105 | 5.69 | 5.54 | 0.01 | 0.01 | 8.42
iPOne013110 | iPOne013 | 110 | 5.5 | 5.35 | 0 | -0.01 | 9.14
iPOne013115 | iPOne013 | 115 | 5.33 | 5.19 | 0.03 | -0.02 | 9.89
iPOne013120 | iPOne013 | 120 | 5.15 | 5.04 | 0.03 | -0.02 | 10.66
iPOne013125 | iPOne013 | 125 | 5.07 | 4.94 | 0.02 | -0.01 | 11.46
iPOne013130 | iPOne013 | 130 | 4.9 | 4.76 | 0.01 | -0.03 | 12.28
iPOne013135 | iPOne013 | 135 | 4.77 | 4.63 | 0.02 | -0.05 | 13.12
iPOne013140 | iPOne013 | 140 | 4.63 | 4.49 | 0.02 | -0.05 | 13.99
iPOne013145 | iPOne013 | 145 | 4.49 | 4.38 | -0.02 | -0.01 | 14.87
iPOne013150 | iPOne013 | 150 | 4.39 | 4.28 | 0.01 | -0.04 | 15.78
iPOne013155 | iPOne013 | 155 | 4.29 | 4.18 | -0.01 | -0.04 | 16.71
iPOne013160 | iPOne013 | 160 | 4.18 | 4.07 | -0.02 | -0.04 | 17.66
iPOne013165 | iPOne013 | 165 | 4.05 | 3.91 | 0 | -0.03 | 18.63
iPOne013170 | iPOne013 | 170 | 3.96 | 3.83 | -0.02 | -0.05 | 19.62
iPOne013175 | iPOne013 | 175 | 3.88 | 3.78 | -0.03 | -0.03 | 20.63
iPOne013180 | iPOne013 | 180 | 3.82 | 3.72 | -0.04 | -0.05 | 21.66
iPOne013185 | iPOne013 | 185 | 3.76 | 3.67 | -0.01 | -0.05 | 22.71
iPOne013190 | iPOne013 | 190 | 3.71 | 3.64 | -0.03 | -0.04 | 23.77
iPOne013195 | iPOne013 | 195 | 3.67 | 3.58 | -0.02 | -0.02 | 24.86
iPOne013200 | iPOne013 | 200 | 3.63 | 3.56 | -0.06 | -0.01 | 25.96
iPOne013205 | iPOne013 | 205 | 3.59 | 3.56 | -0.06 | 0.01 | 27.09
iPOne013210 | iPOne013 | 210 | 3.53 | 3.48 | -0.07 | 0.02 | 28.23
iPOne013215 | iPOne013 | 215 | 3.5 | 3.47 | -0.08 | 0.02 | 29.38
iPOne013220 | iPOne013 | 220 | 3.52 | 3.46 | -0.08 | 0.01 | 30.56
iPOne013226.2 | iPOne013 | 226.2 | 3.53 | 3.49 | -0.1 | 0.02 | 32.04
MachineEnergy (MeV) | Machine | Energy (MeV) | sigmaX (mm) | sigmaY (mm) | skewnessX | skewnessY | Range in TCS | iPOne004 Baseline measurements from 06/02/2018, analysis with spotemkine
iPOne00470 | iPOne004 | 70 | 7.887 | 7.769 | -0.017 | -0.012 | 4.08
iPOne00475 | iPOne004 | 75 | 7.59 | 7.397 | -0.021 | -0.02 | 4.62
iPOne00480 | iPOne004 | 80 | 7.247 | 7.063 | -0.019 | -0.012 | 5.18
iPOne00485 | iPOne004 | 85 | 6.904 | 6.722 | -0.03 | -0.025 | 5.78
iPOne00490 | iPOne004 | 90 | 6.55 | 6.376 | -0.029 | -0.009 | 6.4
iPOne00495 | iPOne004 | 95 | 6.241 | 6.104 | -0.033 | -0.012 | 7.05
iPOne004100 | iPOne004 | 100 | 5.984 | 5.849 | -0.014 | -0.007 | 7.72
iPOne004105 | iPOne004 | 105 | 5.799 | 5.628 | -0.014 | -0.006 | 8.42
iPOne004110 | iPOne004 | 110 | 5.615 | 5.443 | -0.031 | -0.016 | 9.14
iPOne004115 | iPOne004 | 115 | 5.442 | 5.279 | -0.017 | -0.025 | 9.89
iPOne004120 | iPOne004 | 120 | 5.26 | 5.077 | -0.015 | -0.032 | 10.66
iPOne004125 | iPOne004 | 125 | 5.123 | 4.971 | -0.012 | -0.003 | 11.46
iPOne004130 | iPOne004 | 130 | 4.972 | 4.805 | -0.031 | -0.033 | 12.28
iPOne004135 | iPOne004 | 135 | 4.818 | 4.703 | 0 | -0.021 | 13.12
iPOne004140 | iPOne004 | 140 | 4.695 | 4.6 | -0.017 | -0.052 | 13.99
iPOne004145 | iPOne004 | 145 | 4.614 | 4.493 | 0.01 | -0.038 | 14.87
iPOne004150 | iPOne004 | 150 | 4.495 | 4.379 | 0.011 | -0.017 | 15.78
iPOne004155 | iPOne004 | 155 | 4.395 | 4.29 | 0.014 | -0.037 | 16.71
iPOne004160 | iPOne004 | 160 | 4.306 | 4.193 | 0.007 | -0.051 | 17.66
iPOne004165 | iPOne004 | 165 | 4.188 | 4.076 | 0.018 | -0.061 | 18.63
iPOne004170 | iPOne004 | 170 | 4.088 | 3.986 | 0.014 | -0.05 | 19.62
iPOne004175 | iPOne004 | 175 | 4.01 | 3.879 | 0.028 | -0.035 | 20.63
iPOne004180 | iPOne004 | 180 | 3.92 | 3.804 | -0.001 | -0.046 | 21.66
iPOne004185 | iPOne004 | 185 | 3.865 | 3.759 | 0.014 | -0.054 | 22.71
iPOne004190 | iPOne004 | 190 | 3.833 | 3.697 | 0.014 | -0.012 | 23.77
iPOne004195 | iPOne004 | 195 | 3.807 | 3.693 | 0.02 | -0.047 | 24.86
iPOne004200 | iPOne004 | 200 | 3.783 | 3.686 | 0.026 | -0.017 | 25.96
iPOne004205 | iPOne004 | 205 | 3.76 | 3.695 | 0.022 | 0.034 | 27.09
iPOne004210 | iPOne004 | 210 | 3.77 | 3.656 | 0.018 | 0.024 | 28.23
iPOne004215 | iPOne004 | 215 | 3.701 | 3.611 | -0.006 | 0.03 | 29.38
iPOne004220 | iPOne004 | 220 | 3.623 | 3.542 | 0.023 | 0.023 | 30.56
iPOne004226.2 | iPOne004 | 226.2 | 3.657 | 3.531 | -0.011 | 0.023 | 32.04
MachineEnergy (MeV) | Machine | Energy (MeV) | sigmaX (mm) | sigmaY (mm) | skewnessX | skewnessY | Range in TCS | iPOne006 Baseline measurements from 07/02/2019, analysis with spotemkine
iPOne00670.18 | iPOne006 | 70.18 | 7.939 | 7.847 | -0.024 | 0.014 | 4.08
iPOne00675 | iPOne006 | 75 | 7.529 | 7.454 | -0.019 | 0.015 | 4.62
iPOne00680 | iPOne006 | 80 | 7.189 | 7.066 | -0.016 | 0.01 | 5.18
iPOne00685 | iPOne006 | 85 | 6.858 | 6.713 | -0.027 | 0.001 | 5.78
iPOne00690 | iPOne006 | 90 | 6.542 | 6.414 | -0.031 | -0.005 | 6.4
iPOne00695 | iPOne006 | 95 | 6.248 | 6.162 | -0.029 | -0.003 | 7.05
iPOne006100 | iPOne006 | 100 | 6.006 | 5.926 | -0.008 | -0.013 | 7.72
iPOne006105 | iPOne006 | 105 | 5.774 | 5.678 | -0.011 | -0.008 | 8.42
iPOne006110 | iPOne006 | 110 | 5.565 | 5.444 | -0.028 | -0.014 | 9.14
iPOne006115 | iPOne006 | 115 | 5.397 | 5.31 | -0.013 | 0.002 | 9.89
iPOne006120 | iPOne006 | 120 | 5.234 | 5.182 | -0.002 | -0.017 | 10.66
iPOne006125 | iPOne006 | 125 | 5.141 | 5.034 | -0.017 | -0.01 | 11.46
iPOne006130 | iPOne006 | 130 | 4.973 | 4.872 | -0.001 | -0.019 | 12.28
iPOne006135 | iPOne006 | 135 | 4.846 | 4.738 | 0.004 | -0.04 | 13.12
iPOne006140 | iPOne006 | 140 | 4.692 | 4.589 | -0.012 | -0.015 | 13.99
iPOne006145 | iPOne006 | 145 | 4.556 | 4.443 | 0.017 | -0.04 | 14.87
iPOne006150 | iPOne006 | 150 | 4.434 | 4.329 | -0.006 | -0.042 | 15.78
iPOne006155 | iPOne006 | 155 | 4.334 | 4.209 | 0.014 | -0.044 | 16.71
iPOne006160 | iPOne006 | 160 | 4.274 | 4.129 | -0.014 | -0.045 | 17.66
iPOne006165 | iPOne006 | 165 | 4.135 | 4.075 | -0.015 | -0.085 | 18.63
iPOne006170 | iPOne006 | 170 | 4.041 | 3.947 | -0.008 | -0.049 | 19.62
iPOne006175 | iPOne006 | 175 | 3.939 | 3.856 | -0.034 | -0.07 | 20.63
iPOne006180 | iPOne006 | 180 | 3.833 | 3.759 | -0.002 | -0.037 | 21.66
iPOne006185 | iPOne006 | 185 | 3.774 | 3.68 | 0.007 | -0.04 | 22.71
iPOne006190 | iPOne006 | 190 | 3.704 | 3.633 | -0.023 | -0.03 | 23.77
iPOne006195 | iPOne006 | 195 | 3.661 | 3.555 | -0.036 | -0.02 | 24.86
iPOne006200 | iPOne006 | 200 | 3.614 | 3.482 | -0.021 | -0.021 | 25.96
iPOne006205 | iPOne006 | 205 | 3.544 | 3.424 | -0.051 | 0.003 | 27.09
iPOne006210 | iPOne006 | 210 | 3.443 | 3.356 | 0 | -0.009 | 28.23
iPOne006215 | iPOne006 | 215 | 3.401 | 3.326 | -0.012 | 0.017 | 29.38
iPOne006220 | iPOne006 | 220 | 3.404 | 3.261 | -0.019 | 0.029 | 30.56
iPOne006226.09 | iPOne006 | 226.09 | 3.355 | 3.246 | -0.026 | -0.009 | 32.04

## Baseline kV
Baseline kV charge measurements
Baseline values for the currently selected machine:
| Lookup index (Monthly QA sheet): | 1 | 2 | 3
|  | Image Acquisition Type/Preset | Chamber SN | TP corrected baseline (pC)
|  | CBCT "Head" | #N/A | #N/A
|  | CBCT "Pelvis" | #N/A | #N/A
|  | Oblique OB1 "Pelvis" | #N/A | #N/A
|  | Oblique OB2 "Pelvis" | #N/A | #N/A
Individual machine baselines (add new/modify baselines in the table below):
| Lookup index (for above): | 3 | 4 | 5 | 6 | 7 | 8 | 9
Vlookup Key | Machine | Image Acquisition Type/Preset | Chamber SN | TP corrected baseline (pC) | Date | Temperature (°C) | Pressure (hPa) | Charge (pC)
iPOne004CBCT "Head" | iPOne004 | CBCT "Head" | 433 | 171.2179 | 2021-08-05 | 23.5 | 994.1 | 166
iPOne004CBCT "Pelvis" | iPOne004 | CBCT "Pelvis" | 433 | 949.0726 | 2021-08-05 | 25.3 | 994.1 | 914.6
iPOne004Oblique OB1 "Pelvis" | iPOne004 | Oblique OB1 "Pelvis" | 433 | 0.79902 | 2021-08-05 | 25.3 | 994.1 | 0.77
iPOne004Oblique OB2 "Pelvis" | iPOne004 | Oblique OB2 "Pelvis" | 433 | 0.8094 | 2021-08-05 | 25.3 | 994.1 | 0.78
iPOne006CBCT "Head" | iPOne006 | CBCT "Head" |  | 175.6208 | 2022-01-25 | 20.9 | 1025.5 | 177.2
iPOne006CBCT "Pelvis" | iPOne006 | CBCT "Pelvis" |  | 958.8777 | 2022-01-25 | 20.9 | 1025.5 | 967.5
iPOne006Oblique OB1 "Pelvis" | iPOne006 | Oblique OB1 "Pelvis" |  | 0.83251 | 2022-01-25 | 20.9 | 1025.5 | 0.84
iPOne006Oblique OB2 "Pelvis" | iPOne006 | Oblique OB2 "Pelvis" |  | 0.83251 | 2022-01-25 | 20.9 | 1025.5 | 0.84
iPOne013CBCT "Head" | iPOne013 | CBCT "Head" | 451 | 156.1416 | 2022-01-28 | 21.1 | 1025.9 | 157.5
iPOne013CBCT "Pelvis" | iPOne013 | CBCT "Pelvis" | 451 | 882.225 | 2022-01-28 | 21.1 | 1025.9 | 889.9
iPOne013Oblique OB1 "Pelvis" | iPOne013 | Oblique OB1 "Pelvis" | 451 | 0.7931 | 2022-01-28 | 21.1 | 1025.9 | 0.8
iPOne013Oblique OB2 "Pelvis" | iPOne013 | Oblique OB2 "Pelvis" | 451 | 0.77327 | 2022-01-28 | 21.1 | 1025.9 | 0.78

## Version History
Current: | v3.1 10/05/2022
Version | Date | Changes | Changed by
1 | 2018-04-25 | Modified WPE Monthly QA sheet, added new baselines and changed all the descriptions | Jamil Lambert
1.1 | 2018-07-17 | Changed: "8. Proton beam to X-ray isocentre alignment" to use XRV-100 instead of film. Added section "7. Lateral Uniformity" and fixed subsequent numbering | Jamil Lambert
1.2 | 2018-09-27 | Added description to "5. Range and SOBP.." to use Octave analysis and changed the baseline values from Incline analysis to Ocatave analysis | Jamil Lambert
1.3 | 2018-10-10 | Changed "6.Spot size …" position > 1 mm off tolerance from 5% to 1 spot. And to have six energies instead of 3, reorientated results table. Ocave script also changed to match. | Jamil Lambert
1.4 | 2019-04-15 | Modified NP10 to include NE22 baseline measurements | Jamil Lambert
1.5 | 2019-08-22 | Previously modified to include RG2 baseline measurements, locked tabs and saved with new version number today | Edward Ilsley, Jamil Lambert
1.6 | 2019-08-29 | Modified RG2 Baseline Monoenergetic tab dose in water (column D). Values changed from those in water with isocenter at surface (commissioning measurements on 9/5/19) to those in water with isocenter at 2cm depth (measured on 23/8/19). EI and FF confirmed changes to be made were correct with JL. | Edward Ilsley
1.7 | 2019-11-12 | Changed linearity test from a 10x10 to a 5x5 single energy layer | Jamil Lambert
1.8 | 2019-12-17 | Changed 2. Dose-MU linearity "Beam" in description to a 10x10 cm field, was 5x5 | Edward Ilsley
1.9 | 2020-02-11 | 9. corrected label of position error from (mm) to (cm) in table | Jamil Lambert
2 | 2020-01-04 | Sections included for kV 3D image quality, kV 2D image quality and kV outputs. These are now sections 10, 11 and 12 respectively. Interlocks, daily QA review and comments sections have been relabelled as sections 13, 14 and 15 respectively. | Edward Ilsley
2.1 | 2021-11-09 | Updated Electrometer and Chamber calibration factor table, added check to absolute dose that the calibration is in date | Jamil Lambert
3 | 2022-04-11 | Combined all sites into one excel sheet, based on a modified RG2 with imaging template Added an action level on top of the existing tolerance level. Rounded calculated results to the displayed precision so that values on the tolerance levels behave as expected, e.g. previously 1.0 could pass or fail a 1.0 tolerance. | Jamil Lambert
3.1 | 2022-05-10 | Added the max chamber deviation results to 7. Lateral uniformity | Jamil Lambert